Standard Operating Procedures  

Routine 1H, 13C NMR Data Acquisition
For NSC Bruker 400 NMR (AvanceCore)
Dr. Zhenming Du

Email: zdu@gsu.edu  Office Phone: 3-5538 Office: NSC 134A 
Revised 10/28/2025
LC=Left Click; DC=Double left Click.
Black and bold command need to be typed and followed with <etr>.
NSC 400 Data Directory: C:\Bruker\Topspin4.5.0\data\‘userid’\nmr
PSC 400 Data Directory: C:\Bruker\Topspin3.7pl6\data\nmr\‘userid’
	Step
	Command
	Notes

	1
	Login 
	DO NOT share your user ID with unauthorized person. 

	2
	Start Topspin program by DC the icon. 
	NSC: Use Topspin4.5.0 (Subject to change).
PSC: Use Topspin3.1 (Subject to change).

	3
	Open  one of your old NMR file, in command line, type new and <etr>
Thus your data will be in your directory.
	In the new dataset window: 
Change the Name  or Expno or both. 
1) Select to use the current parameters;

2) Do not worry about the solvent;

3) Additional actions choose no.

You always need a new file for NMR data. 
Do not use ‘-’ or  ‘space’ in naming. 

	4
	bsmsdisp  <etr>

	Make sure on-off button below LOCK is not highlighted;
Make sure spin button below SAMPLE is not highlighted. 

	5
	Measure the depth of your sample and load it to position XX of the carousel. . Carousel rotate clockwise only. 
	XX stands for the position of the sample. It corresponds to a number between 1-24. 

Only use nmr tube that are good!

Broken NMR Tubes are prohibited!
Sample volume need to be 4 cm or higher!

	6
	sx XX <etr>
XX is between 1-24.

It is the position of your sample from step 5.
	This will Eject sample from the magnet if there is one inside and put your sample in.
Walk carefully around the magnet.
Double check that the sample status on the bsmsdisp window showing ‘down’ after execution of command. 

Close BSMS Control Suite Window when done.

	7
	rpar <etr>
	Read the appropriate parameter set for various expt.

For 1-D proton NMR, rpar ‘Proton’;

For 1-D Carbon NMR, rpar “C13CPD’;

	8
	lockdisp <etr>


	Lock display window opens.


	9
	rsh <etr>
	Shimming file lists opens. LC on the file “lastbestbbo” or some similar files.

	10
	lock <etr>


	LC on the appropriate lock solvent. Wait until “lockn:done” message appears.

	11
	atma <etr>
	Automatic probe tuning and match starts. 

Wait till “atma:finished” message appears. For 13C experiments, please wait till both carbon and proton channels tuning are done.

	12
	topshim 2h  <etr>
	Automatic shimming starts. 

Wait until “topshim completed” message appears. 

	13 (in case step 12 not working)
	bsmsdisp  <etr>
	Open BSMS board for manual shimming. Tune on Z, Z2,and Z with size of 5, 10, and 5 to maximize the lock signal. Adjust gain under Auto section if necessary. Close BSMS Control Suite Window.  

	14
	ased <etr>
	AquProc setup mode.  

	15
	getprosol <etr>
	Read corresponding probe parameters.

	16
	rga <etr>


	Automatic receiver gain optimization starts. 

Wait for the “rga:finished” message to appear.

	17
	expt <etr>
	Calculate the time to finish the experiment;

Adjust “ns” if the acquisition time is too long;

DO NOT EAT other user’s time.

	18
	zg <etr>


	Start data acquisition.
Wait till the “acquisition: finished” message to appear. 

	19
	efp <etr>

apk <etr>

abs <etr>

Command ‘duzproc’ does all three commands together.
	Type ‘duzproc’ for simple data processing.

If doing 13C, type ‘lb 5’ before processing.
If additional experiments are required, repeat steps 3, 7, (11?),14,15,16,17,18. 
Type tr to transfer data for processing if exp not finished;

Type halt to stop experiment and save data;

STOP ALL UNFINISHED EXP BEFORE STEP 20

	20
	bsmsdisp <etr>
	Make sure on-off button below LOCK is not highlighted; Make sure spin button below SAMPLE is not highlighted.

	21
	Load D2O sample to position xx.
	Calibrate D2O sample depth.


	22
	sx xx  <etr>
	the D2O sample is inserted into the magnet. Wait until the sample status on the bsmsdisp window button showing down. Then close BSMSDISP display window.

	23
	lock d2o <etr>


	Locks d2o sample. Wait until “lockn:finished” message appears.

	24
	Transfer your data;

Back up your data.
	Please process this within 2 weeks’ timeframe as old data may be deleted without notice. Backup for satefy.

	25
	exit <etr> 

	Always type ‘exit’ to terminate the program. Please DO NOT click the crossbar as it does not always terminate all active commands.
Log out PC. 
Fill out the log book 

Record all error messages from the PC screen to logbook if instruments are malfunctioning. Email or call Dr.Du if the status is NOT ok.




TOPSPIN Processing Guide

Dummy Data Processing guide

tr: transfer unfinished 1D fid;

efp: em+ft+ph

apk: automatic phase correction

abs: automatic baseline correction

Detailed Data Processing guide
1.   Opening Files and FTing Data

a.   Click on File -> Open and find your file name.

b.   Click OK.  If multiple experiments were run with that filename, you may have to select an EXPNO.

c.   Type “efp” to FT your data.

2.   Zooming In / Zooming Out

a.   [image: image1.png]2 j2 *8 /8



 Increase/decrease vertical scale by factor of either 2 or 8. b.   [image: image2.png]


 Press LMB down while dragging mouse up/down to inc/dec vertical

scale manually.

c.   [image: image3.png]


 Scales to tallest peak.

d.   [image: image4.png]


 Scales out full sweep width.

e.   [image: image5.png]


 Moves spectrum upfield/downfield. f.
[image: image6.png]


 Moves spectrum up/down.

g.   [image: image7.png]


 Switches scale between Hz and ppm.

3.   Zooming In / Out Using the Mouse

4.   Manual Phasing

a.   Click on [image: image8.png]


 to enter phase mode.

b.   The buttons above your spectrum should look like this:

[image: image9.png]Ao 1 R|0 ww0|a e 0B E I




c.   Click and hold the LMB on the [image: image10.png]


 button to perform zero-order phasing. d.   Click and hold the LMB on the [image: image11.png]


 button to perform first-order

phasing.

e.   Click on [image: image12.png]


 to save and return when you are done phasing.

5.   Calibrating the Spectrum a.   Click on [image: image13.png]


.

b.   Click on your reference peak (solvent or TMS). c.   Enter the chemical shift of that peak.

6.   Integration

a.   NOTE:  It is easiest to zoom in on a small region and use the [image: image14.png]


 buttons to move around while integrating!

b.   Click on the [image: image15.png]


 button to enter integrals mode.  The toolbar above your spectrum should change to this:

[image: image16.png]T d Fola@|hsf #|lan n|x[=|2p2




c.   Click on [image: image17.png]


.  The box should turn green, indicating you are now integrating.

d.   Click and hold with the LMB on the left side of the peak you want to integrate, and drag while holding down LMB to the right side, then let go when you are done integrating.

e.   Repeat step d for each peak or region of interest.

f.
To adjust bias and slope, click and hold LMB on [image: image18.png]s



 while dragging the mouse.

g.   To Open/Save a region file:

1.   Click on [image: image19.png]an



 to open or save a region file already associated with that spectrum.

h.   To Delete an integral:

1.   Right click on the region you want to delete and choose “Select”.

Once selected, the integral region should turn yellow.

2.   You can select many regions at once by repeating step 1.

3.   Click on [image: image20.png]


 to delete the selected regions.

i.
When you are done, click [image: image21.png]


 to save and return the integral file.

7.   Peak Picking

a.   Click on [image: image22.png]


 to enter peak picking mode.  The toolbar above your spectrum should change to look like this:

[image: image23.png]



b.   Place your cursor in the upper left hand corner of the spectrum.

c.   Click and drag LMB down to the lower right hand corner of the spectrum. d.   The maximum and minimum are automatically set based on the region

you select.

e.   To delete your peak picking regions, click [image: image24.png]


. f.
Click on [image: image25.png]


 to save and return.

8.   Manual Peak Picking

a.   Click on [image: image26.png]


 (green) to manually select peaks.  A red cursor will appear. b.   Click the LMB on each peak you want to pick.

c.   Click [image: image27.png]


to save and return.

9.   Plotting

a.   Type ‘plot’.

b.   A default layout of the 1-D spectrum appears.

c.   To adjust the horizontal scale, click the spectrum area.

1.   Select ‘Edit’.

2.   Change the x-axis to be whatever range you need (ie 10 – 0 ppm)

3.   Click Apply and OK when you are done;

d.   To adjust the vertical scale, move the crosshair into the spectrum area.

1.   Select ‘1D/2D-Edit’.

2.   Use the buttons to adjust the vertical scale of the spectrum.

3.   Click Apply and OK.

e.   Click on ‘File’.

f.
Click on ‘Print’.

